## Slide 1: The Algorithm A in Detail S. Subramanian

## Slide 2
Where to find details?
ISO 5725-5:1998, clause 6.2
ISO 13528:2015, ANNEXURE E3

## Slide 3: Algorithm A – Action 1
Arrange the "p" lab results in the increasing order. |  | Lab | Value |  | Lab | Value
 |  | 1 | 1.69 |  | 2 | 0.74
 |  | 2 | 0.74 |  | 20 | 0.80
 |  | 6 | 2.05 |  | 8 | 1.14
 |  | 8 | 1.14 |  | 22 | 1.21
 |  | 9 | 2.19 |  | 12 | 1.39
 |  | 12 | 1.39 |  | 14 | 1.50
 |  | 13 | 1.52 |  | 13 | 1.52
 |  | 14 | 1.50 |  | 19 | 1.58
 |  | 19 | 1.58 |  | 25 | 1.63
 |  | 20 | 0.80 |  | 1 | 1.69
 |  | 22 | 1.21 |  | 6 | 2.05
 |  | 25 | 1.63 |  | 9 | 2.19

## Slide 4: Algorithm A – Action 2
Calculate the median of "p" lab results as Initial X*. 2)Calculate the absolute of the deviation of each lab results from "Initial X* 3) Arrange the "p" deviations in the increasing order. 4) Calculate the median of "p" deviations. 5) Multiply the median of deviations by 1.483 (constant) to get "Initial s*. |  | Lab | Value | Deviation |  | Deviation
 |  | 2 | 0.74 | 0.77 |  | 0.01
 |  | 20 | 0.80 | 0.71 |  | 0.01
 |  | 8 | 1.14 | 0.37 |  | 0.07
 |  | 22 | 1.21 | 0.30 |  | 0.12
 |  | 12 | 1.39 | 0.12 |  | 0.12
 |  | 14 | 1.50 | 0.01 |  | 0.18
 |  | 13 | 1.52 | 0.01 |  | 0.30
 |  | 19 | 1.58 | 0.07 |  | 0.37
 |  | 25 | 1.63 | 0.12 |  | 0.54
 |  | 1 | 1.69 | 0.18 |  | 0.68
 |  | 6 | 2.05 | 0.54 |  | 0.71
 |  | 9 | 2.19 | 0.68 |  | 0.77
 |  | Initial X* | 1.510 |  | Median= | 0.240
 |  | Initial S* | 0.3559 |  |  | 

## Slide 5
We now have for the set of “p” results
Initial X*= 1.510
Initial s*= 0.3559
For iteration 1, calculate d as
d = 1.5 x s* = 1.5 x 0.3559 = 0.5339
Calculate X* - d = 1.510 –0.5339 = 0.9761 =0.98
X*+ d = 1.510+0.5339 = 2.0439 = 2.04
Compare the original “p” results with 0.98 (minimum permitted result) and 2.04 (maximum permitted result). If any result is < 0.98 change it as 0.98. If any result is > 2.04 change it as 2.04. If results are between 0.98 and 2.04 do not change them.
We now have a new set of “p” values for iteration 1.

## Slide 6
For the new set of “p” results of Iteration 1, calculate
New X* = Average of the new set of results = 1.475
New s*= 1.134 x SD of new set of results = 0.4072
For iteration 2, calculate d as
d = 1.5 x s* = 1.5 x 0.4072= 0.6109
Calculate X* - d = 1.475 – 0.6109= 0.8641 =0.86
X*+ d= 1.475 + 0.6109= 2.0859 = 2.09
Compare the original “p” results with 0.86 (minimum permitted result) and 2.09 (maximum permitted result). If any result is < 0.86 change it as 0.86. If any result is > 2.09 change it as 2.09. If results are between 0.86 and 2.09 do not change them.
We now have a new set of “p” values for iteration 2.

## Slide 7
For the new set of “p” results of Iteration 2, calculate
New X* = Average of the new set of results = 1.460
New s*= 1.134 x SD of new set of results = 0.4486
For iteration 3, calculate d as
d = 1.5 x s* = 1.5 x 0.4486= 0.6728
Calculate X* - d = 1.460 – 0.6728= 0.7875=0.79
X*+ d= 1.460 + 0.6728= 2.1332 = 2.13
Compare the original “p” results with 0.79 (minimum permitted result) and 2.13 (maximum permitted result). If any result is < 0.79 change it as 0.79. If any result is > 2.13 change it as 2.13. If results are between 0.79 and 2.13 do not change them.
We now have a new set of “p” values for iteration 3.

## Slide 8
For the new set of “p” results of Iteration 3, calculate
New X* = Average of the new set of results = 1.453
New s*= 1.134 x SD of new set of results = 0.4786
Repeat the above process for Iteration 4, 5, 6… and get New X* and New s*, until the New X* and New s* are converging to a minimum of 3 significant figures.
Treat the latest “New X*” value as Assigned Value, xpt and “New s*” value as SDPA σpt.
Calculate std. uncertainty in Assigned Value, u(xpt) as follows: u(xpt) = 1.25/ √p x s*

## Slide 9: Algorithm A – Action 3
1) Copy the original values of "p" results for 5 iterations. 2) For iteration 1, calculate d= 1.5 x S*. 3) Calculate X*- d and X* + d for iteration 1 4) If any result of iteration 1 is below X*- d change it as X*- d 5) If any result of iteration 1 is above X*+ d change it as X*+ d 6) If the result is between X*- d and X*+ d retain it 7) For new set of iteration 1 calculate average as "new X*" 8) For new set of iteration 1 calculate 1.134 x SD of the new set as "new s*" 9) repeat the process fro step 2 for iteration 2 to 8 10) Continue this till both "new x*" and :New s*" converge to 3 significant figures. |  | d =1.5 x S* |  |  | 0.53 |  |  |  |  |  |  |  | 
 |  | x* - d |  |  | 0.98 |  |  |  |  |  |  |  | 
 |  | x* + d |  |  | 2.04 |  |  |  |  |  |  |  | 
 |  | Lab | Value | Deviation | 1 |  | 2 |  | 3 |  | 4 |  | 5
 |  | 2 | 0.74 | 0.77 | 0.98 |  | 0.74 |  | 0.74 |  | 0.74 |  | 0.74
 |  | 20 | 0.80 | 0.71 | 0.98 |  | 0.80 |  | 0.80 |  | 0.80 |  | 0.80
 |  | 8 | 1.14 | 0.37 | 1.14 |  | 1.14 |  | 1.14 |  | 1.14 |  | 1.14
 |  | 22 | 1.21 | 0.30 | 1.21 |  | 1.21 |  | 1.21 |  | 1.21 |  | 1.21
 |  | 12 | 1.39 | 0.12 | 1.39 |  | 1.39 |  | 1.39 |  | 1.39 |  | 1.39
 |  | 14 | 1.50 | 0.01 | 1.50 |  | 1.50 |  | 1.50 |  | 1.50 |  | 1.50
 |  | 13 | 1.52 | 0.01 | 1.52 |  | 1.52 |  | 1.52 |  | 1.52 |  | 1.52
 |  | 19 | 1.58 | 0.07 | 1.58 |  | 1.58 |  | 1.58 |  | 1.58 |  | 1.58
 |  | 25 | 1.63 | 0.12 | 1.63 |  | 1.63 |  | 1.63 |  | 1.63 |  | 1.63
 |  | 1 | 1.69 | 0.18 | 1.69 |  | 1.69 |  | 1.69 |  | 1.69 |  | 1.69
 |  | 6 | 2.05 | 0.54 | 2.04 |  | 2.05 |  | 2.05 |  | 2.05 |  | 2.05
 |  | 9 | 2.19 | 0.68 | 2.04 |  | 2.19 |  | 2.19 |  | 2.19 |  | 2.19
 |  | Initial X* | 1.510 | 0.240 | 1.475 |  |  |  |  |  |  |  | 
 |  | Initial S* | 0.3559 |  | 0.4072 |  |  |  |  |  |  |  | 

## Slide 10: Algorithm A – Action 4
Process is repeated till Iteration 6. While "new x*" has converged upto 4 significant figures ( 3 decimal places) "new s*" has converged upto 5 significant figures (4 decimal places) at iteration 6. |  | d =1.5 x S* |  |  | 0.53 |  | 0.61 |  | 0.67 |  | 0.72 |  | 0.74 |  | 0.74
 |  | x* - d |  |  | 0.98 |  | 0.86 |  | 0.79 |  | 0.73 |  | 0.71 |  | 0.71
 |  | x* + d |  |  | 2.04 |  | 2.09 |  | 2.13 |  | 2.17 |  | 2.19 |  | 2.20
 |  | Lab | Value | Deviation | 1 |  | 2 |  | 3 |  | 4 |  | 5 |  | 6
 |  | 2 | 0.74 | 0.77 | 0.98 |  | 0.86 |  | 0.79 |  | 0.74 |  | 0.74 |  | 0.74
 |  | 20 | 0.80 | 0.71 | 0.98 |  | 0.86 |  | 0.80 |  | 0.80 |  | 0.80 |  | 0.80
 |  | 8 | 1.14 | 0.37 | 1.14 |  | 1.14 |  | 1.14 |  | 1.14 |  | 1.14 |  | 1.14
 |  | 22 | 1.21 | 0.30 | 1.21 |  | 1.21 |  | 1.21 |  | 1.21 |  | 1.21 |  | 1.21
 |  | 12 | 1.39 | 0.12 | 1.39 |  | 1.39 |  | 1.39 |  | 1.39 |  | 1.39 |  | 1.39
 |  | 14 | 1.50 | 0.01 | 1.50 |  | 1.50 |  | 1.50 |  | 1.50 |  | 1.50 |  | 1.50
 |  | 13 | 1.52 | 0.01 | 1.52 |  | 1.52 |  | 1.52 |  | 1.52 |  | 1.52 |  | 1.52
 |  | 19 | 1.58 | 0.07 | 1.58 |  | 1.58 |  | 1.58 |  | 1.58 |  | 1.58 |  | 1.58
 |  | 25 | 1.63 | 0.12 | 1.63 |  | 1.63 |  | 1.63 |  | 1.63 |  | 1.63 |  | 1.63
 |  | 1 | 1.69 | 0.18 | 1.69 |  | 1.69 |  | 1.69 |  | 1.69 |  | 1.69 |  | 1.69
 |  | 6 | 2.05 | 0.54 | 2.04 |  | 2.05 |  | 2.05 |  | 2.05 |  | 2.05 |  | 2.05
 |  | 9 | 2.19 | 0.68 | 2.04 |  | 2.09 |  | 2.13 |  | 2.17 |  | 2.19 |  | 2.19
 |  | Initial X* | 1.510 | 0.240 | 1.475 |  | 1.460 |  | 1.453 |  | 1.452 |  | 1.453 |  | 1.453
 |  | Initial S* | 0.3559 |  | 0.4072 |  | 0.4486 |  | 0.4786 |  | 0.4928 |  | 0.4961 |  | 0.4961