## Slide 1: How to do Grubbs’ test S.SUBRAMANIAN

## Slide 2: How to do Grubbs’ test?
Arrange the participants results in the increasing order
Calculate Average and SD for this set of data
Calculate Gmax= (xmax – x)/SD, where Xmax is the maximum result reported and X is the average
Compare G with G critical value at 95% CL
If G > Gcrit conclude that it is outlier and remove this data from the data set

## Slide 3: How to do Grubbs’ test?
Calculate Gmin = (x - xmin)/SD, where Xmin is the minimum result reported and X is the average
Compare G with G critical value at 95% CL
If G > Gcrit conclude that it is outlier and remove this data from the data set
9) Repeat the above till both max and min values of the new set are not outliers

## Slide 4: How to do Grubbs’ test?
Use the new set after removal of outliers for calculating both assigned value (and its uncertainty ) + SDPA
11) The performance evaluation of the removed results has also to be conducted using the assigned value & SDPA determined as specified above

## Slide 5: Gcritical values at 95% CL
No. of data set (n) | Gcrit | No. of data set (n) | Gcrit | No. of data set (n) | Gcrit | No. of data set (n) | Gcrit | No. of data set (n) | Gcrit
-- | -- | 21 | 2.733 | 41 | 3.050 | 61 | 3.206 | 81 | 3.314
-- | -- | 22 | 2.758 | 42 | 3.060 | 62 | 3.212 | 82 | 3.318
3 | 1.155 | 23 | 2.781 | 43 | 3.070 | 63 | 3.218 | 83 | 3.322
4 | 1.481 | 24 | 2.802 | 44 | 3.080 | 64 | 3.224 | 84 | 3.326
5 | 1.715 | 25 | 2.822 | 45 | 3.090 | 65 | 3.230 | 85 | 3.330
6 | 1.887 | 26 | 2.841 | 46 | 3.098 | 66 | 3.236 | 86 | 3.334
7 | 2.020 | 27 | 2.859 | 47 | 3.106 | 67 | 3.242 | 87 | 3.338
8 | 2.126 | 28 | 2.876 | 48 | 3.114 | 68 | 3.248 | 88 | 3.342
9 | 2.215 | 29 | 2.893 | 49 | 3.122 | 69 | 3.254 | 89 | 3.346
10 | 2.290 | 30 | 2.908 | 50 | 3.130 | 70 | 3.260 | 90 | 3.350
11 | 2.355 | 31 | 2.924 | 51 | 3.137 | 71 | 3.265 | 91 | 3.353
12 | 2.412 | 32 | 2.938 | 52 | 3.144 | 72 | 3.270 | 92 | 3.356
13 | 2.462 | 33 | 2.952 | 53 | 3.151 | 73 | 3.275 | 93 | 3.359
14 | 2.507 | 34 | 2.965 | 54 | 3.158 | 74 | 3.280 | 94 | 3.362
15 | 2.549 | 35 | 2.979 | 55 | 3.165 | 75 | 3.285 | 95 | 3.365
16 | 2.585 | 36 | 2.991 | 56 | 3.172 | 76 | 3.290 | 96 | 3.368
17 | 2.620 | 37 | 3.003 | 57 | 3.179 | 77 | 3.295 | 97 | 3.371
18 | 2.651 | 38 | 3.014 | 58 | 3.186 | 78 | 3.300 | 98 | 3.374
19 | 2.681 | 39 | 3.025 | 59 | 3.193 | 79 | 3.305 | 99 | 3.377
20 | 2.709 | 40 | 3.036 | 60 | 3.200 | 80 | 3.310 | 100 | 3.380

## Slide 6: Example of Grubbs test when p ≥12
S.No | Value | G | G crit |  |  |  |  |  |  |  |  |  | 
1 | 59.5 | 2.702 | 2.549 |  | S.No | Value | G | G crit |  |  |  |  | 
2 | 62.4 |  |  |  | 2 | 62.4 | 2.896 | 2.507 |  | S.No | Value | G | G crit
3 | 67.0 |  |  |  | 3 | 67.0 |  |  |  | 3 | 67.0 | 1.792 | 2.462
4 | 67.5 |  |  |  | 4 | 67.5 |  |  |  | 4 | 67.5 |  | 
5 | 67.5 |  |  |  | 5 | 67.5 |  |  |  | 5 | 67.5 |  | 
6 | 68.5 |  |  |  | 6 | 68.5 |  |  |  | 6 | 68.5 |  | 
7 | 69.1 |  |  |  | 7 | 69.1 |  |  |  | 7 | 69.1 |  | 
8 | 69.5 |  |  |  | 8 | 69.5 |  |  |  | 8 | 69.5 |  | 
9 | 69.6 |  |  |  | 9 | 69.6 |  |  |  | 9 | 69.6 |  | 
10 | 70.0 |  |  |  | 10 | 70.0 |  |  |  | 10 | 70.0 |  | 
11 | 70.2 |  |  |  | 11 | 70.2 |  |  |  | 11 | 70.2 |  | 
12 | 70.2 |  |  |  | 12 | 70.2 |  |  |  | 12 | 70.2 |  | 
13 | 70.3 |  |  |  | 13 | 70.3 |  |  |  | 13 | 70.3 |  | 
14 | 70.4 |  |  |  | 14 | 70.4 |  |  |  | 14 | 70.4 |  | 
15 | 70.9 | 0.836 | 2.549 |  | 15 | 70.9 | 0.933 | 2.507 |  | 15 | 70.9 | 1.236 | 2.462
AVG | 68.17 |  |  |  | AVG | 68.79 |  |  |  | AVG | 69.28 |  | 
SD | 3.208 |  |  |  | SD | 2.212 |  |  |  | SD | 1.272 |  | 

## Slide 7: How to do Grubbs’ test? When p<12
Use same method as described earlier, but use
G critical (Gcrit) at 99% CL

## Slide 8: Gcritical values at 99% CL
No. of data set (n) | Gcrit
3 | 1.155
4 | 1.496
5 | 1.764
6 | 1.973
7 | 2.139
8 | 2.274
9 | 2.387
10 | 2.482
11 | 2.564
12 | 2.636

## Slide 9: Example of Grubbs test when p <12
Code | Value | G | Gcrit |  |  |  |  | 
LAB_19 | 16.56 | 2.643 | 2.482 |  | Code | Value | G | Gcrit
LAB_29 | 48.24 |  |  |  | LAB_29 | 48.24 | 1.399 | 2.387
LAB_23 | 49.52 |  |  |  | LAB_23 | 49.52 |  | 
LAB_12 | 51.07 |  |  |  | LAB_12 | 51.07 |  | 
LAB_01 | 54.10 |  |  |  | LAB_01 | 54.10 |  | 
LAB_26 | 55.80 |  |  |  | LAB_26 | 55.80 |  | 
LAB_07 | 57.37 |  |  |  | LAB_07 | 57.37 |  | 
LAB_04 | 59.53 |  |  |  | LAB_04 | 59.53 |  | 
LAB_10 | 61.85 |  |  |  | LAB_10 | 61.85 |  | 
LAB_16 | 62.44 | 0.813 | 2.482 |  | LAB_16 | 62.44 | 1.319 | 2.387
AVG | 51.648 |  |  |  | AVG | 55.547 |  | 
SD | 13.2762 |  |  |  | SD | 5.2243 |  |