## Slide 1: Identification of Blunders in PT results S.SUBRAMANIAN

## Slide 2: Identification of Blunders
Determination of Assigned Value and its uncertainty as stipulated in chapter 7 of ISO 13528:2022:
By formulation (also called as Known Value Scheme);
By using a CRM as PT item;
By characterizing the PT item using a valid CRM in one laboratory (also called as Value transfer from a CRM to a closely matched candidate RM);
By using consensus value from expert laboratories (which are not participants in the PT Scheme); and
By using consensus value from participant laboratories
While methods at (a) to (d) are independent of participants (e) is based on participants’ results only. Any odd value (blunder) present in participants’ results can affect determination of Assigned Value and its uncertainty.

## Slide 3: Identification of Blunders
Determination of SDPA as stipulated in chapter 8 of ISO 13528:2022:
By perception of experts;
By experience from previous rounds of a PT Scheme ;
By use of a General Model (Horwitz equation);
By using repeatability & reproducibility SDs from a previous collaborative study of precision of a measurement method;
From data obtained in the same round of a PT Scheme
While methods at (a) to (d) are independent of participants (e) is based on participants’ results only. Any odd value (blunder) present in participants’ results can affect determination of SDPA.

## Slide 4: Identification of Blunders
Thus, when either Assigned Value, xpt and its uncertainty, u(xpt) and/or SDPA, σpt is to be determined using participants’ results only, then prior to conduct of any statistical analysis of the data, odd values (blunders) have to be identified and removed for the calculation of xpt, u(xpt) as well as σpt .

## Slide 5: How to identify “Blunders”
We will now identify “outliers” (blunders) in the participants’ results by using the following three methods:
Grubbs’ test
Using median & MPE
Results with Z/Z’ score 5 and below or 5 and above

## Slide 6: How to identify “Blunders”
Initially do Grubbs test to identify “outliers” in the participants’ result and remove them for calculating either AV and/or SDPA.
Find the median of the remaining participants’ results . Identify MPE for this parameter in a realistic manner. This is critical.
Calculate Median- MPE and Median + MPE
Exclude participants’ results which are beyond these two limits.
Using the remaining participants’ results calculate AV, SDPA and SU by Algorithm A or nIQR method, as per statistical design selected
Calculate Z score or Z prime score as applicable
Remove those participants results which are – 5 and below as well as +5 and above

## Slide 7: How to identify “Blunders”
8) Using the remaining participants’ results calculate AV, SDPA and SU by Algorithm A or nIQR method, as per statistical design selected
9) Calculate the Z score or Z prime score as applicable for the remaining participants’ results
10) Also calculate Z score or Z prime score as applicable for all the results which are removed for calculating AV, SDPA and SU
11) The Z/Z’ score for the removed results should be in the range of -2 to 2. If it is not so do not consider these results as “outliers” and include them for calculating AV, SDPA and SU.
12) Calculate the Z score or Z prime score as applicable for all the participants’ results including the “outliers”

## Slide 8: Identify Blunder: ORIGINAL RESULTS
Original results |  |  | Results re- arranged | 
Code | Result |  | Code | Result
2 | 314.7 |  | 39 | 248.9
5 | 317.0 |  | 24 | 302.9
9 | 327.8 |  | 10 | 313.8
10 | 313.8 |  | 2 | 314.7
17 | 335.7 |  | 5 | 317.0
19 | 330.5 |  | 32 | 320.0
21 | 328.8 |  | 33 | 322.0
24 | 302.9 |  | 9 | 327.8
27 | 385.1 |  | 21 | 328.8
28 | 339.6 |  | 19 | 330.5
30 | 342.1 |  | 17 | 335.7
32 | 320.0 |  | 37 | 338.4
33 | 322.0 |  | 28 | 339.6
37 | 338.4 |  | 30 | 342.1
39 | 248.9 |  | 27 | 385.1

## Slide 9: Identify Blunder: Grubbs test -GSM
Code | Result | G | Gcrit |  |  |  |  |  |  |  |  |  | 
39 | 248.9 | 2.688 | 2.549 |  | Code | Result | G | Gcrit |  |  |  |  | 
24 | 302.9 |  |  |  | 24 | 302.9 | 1.383 | 2.507 |  | Code | Result | G | Gcrit
10 | 313.8 |  |  |  | 10 | 313.8 |  |  |  | 24 | 302.9 | 1.930 | 2.462
2 | 314.7 |  |  |  | 2 | 314.7 |  |  |  | 10 | 313.8 |  | 
5 | 317.0 |  |  |  | 5 | 317.0 |  |  |  | 2 | 314.7 |  | 
32 | 320.0 |  |  |  | 32 | 320.0 |  |  |  | 5 | 317.0 |  | 
33 | 322.0 |  |  |  | 33 | 322.0 |  |  |  | 32 | 320.0 |  | 
9 | 327.8 |  |  |  | 9 | 327.8 |  |  |  | 33 | 322.0 |  | 
21 | 328.8 |  |  |  | 21 | 328.8 |  |  |  | 9 | 327.8 |  | 
19 | 330.5 |  |  |  | 19 | 330.5 |  |  |  | 21 | 328.8 |  | 
17 | 335.7 |  |  |  | 17 | 335.7 |  |  |  | 19 | 330.5 |  | 
37 | 338.4 |  |  |  | 37 | 338.4 |  |  |  | 17 | 335.7 |  | 
28 | 339.6 |  |  |  | 28 | 339.6 |  |  |  | 37 | 338.4 |  | 
30 | 342.1 |  |  |  | 30 | 342.1 |  |  |  | 28 | 339.6 |  | 
27 | 385.1 | 2.155 | 2.549 |  | 27 | 385.1 | 2.830 | 2.507 |  | 30 | 342.1 | 1.397 | 2.462
Average | 324.49 |  |  |  | Average | 329.89 |  |  |  | Average | 325.64 |  | 
SD | 28.119 |  |  |  | SD | 19.509 |  |  |  | SD | 11.778 |  | 

## Slide 10: Identify Blunder: MPE-GSM
Code | Result |  | 
24 | 302.9 |  | 
10 | 313.8 |  | 
2 | 314.7 |  | 
5 | 317.0 |  | 
32 | 320.0 |  | 
33 | 322.0 |  | 
9 | 327.8 |  | 
21 | 328.8 |  | 
19 | 330.5 |  | 
17 | 335.7 |  | 
37 | 338.4 |  | 
28 | 339.6 |  | 
30 | 342.1 |  | 
Median | 327.80 |  | 
MPE | 19.668 | @6% | 
Median - MPE= |  | 308.1 | 
Median + MPE= |  | 347.5 | 

## Slide 11: Identify Blunder: IzI or Iz’I >5-GSM
Code | Result | z' score
10 | 313.8 | -1.14
2 | 314.7 | -1.06
5 | 317.0 | -0.87
32 | 320.0 | -0.62
33 | 322.0 | -0.46
9 | 327.8 | 0.02
21 | 328.8 | 0.10
19 | 330.5 | 0.25
17 | 335.7 | 0.68
37 | 338.4 | 0.90
28 | 339.6 | 1.00
30 | 342.1 | 1.21
AV= | 327.53 | 
SDPA= | 11.364 | 
SU= | 4.101 | 
No result has Z' beyond 5 |  | 

## Slide 12: Performance evaluation for all results
 | AV= | 327.533
 | SDPA= | 11.3640
 | SU= | 4.1006
Code | Result | z' score
2 | 314.7 | -1.06
5 | 317.0 | -0.87
9 | 327.8 | 0.02
10 | 313.8 | -1.14
17 | 335.7 | 0.68
19 | 330.5 | 0.25
21 | 328.8 | 0.10
24 | 302.9 | -2.04
27 | 385.1 | 4.77
28 | 339.6 | 1.00
30 | 342.1 | 1.21
32 | 320.0 | -0.62
33 | 322.0 | -0.46
37 | 338.4 | 0.90
39 | 248.9 | -6.51
24,27 &39 are excluded. But Z' score of 24 is -2.04 & hence can't be treated as "outlier“. Include code “24” in main group |  | 

## Slide 13: Revised Performance evaluation for all results
 | AV= | 325.951
 | SDPA= | 12.6584
 | SU= | 4.3885
Code | Result | z' score
2 | 314.7 | -0.84
5 | 317.0 | -0.67
9 | 327.8 | 0.14
10 | 313.8 | -0.91
17 | 335.7 | 0.73
19 | 330.5 | 0.34
21 | 328.8 | 0.21
24 | 302.9 | -1.72
27 | 385.1 | 4.41
28 | 339.6 | 1.02
30 | 342.1 | 1.21
32 | 320.0 | -0.44
33 | 322.0 | -0.29
37 | 338.4 | 0.93
39 | 248.9 | -5.75
27 & 39 are only removed for calculation. Z' score of both are >3.0 |  |