## Slide 1: Validation of Assigned Value S.SUBRAMANIAN

## Slide 2: What is validation of Assigned Value?
Assigned Value and its uncertainty can be determined as stipulated in chapter 7 of ISO 13528:2022 by the following 5 methods:
By formulation (also called as Known Value Scheme);
By using a CRM as PT item;
By characterizing the PT item using a valid CRM in one laboratory (also called as Value transfer from a CRM to a closely matched candidate RM);
By using consensus value from expert laboratories (which are not participants in the PT Scheme); and
By using consensus value from participant laboratories
As stated in clause 7.1.2 of ISO 13528:2022, regardless of the method used to determine Assigned Value, validity of the same should be checked for each round of PT Scheme.

## Slide 3: What is validation of Assigned Value?
This means that the assigned value determined by any of the 5 suggested methods should be checked with alternative method as discussed in clause 7.8 of ISO 13528:2022.
Two scenarios can arise:
Assigned value is determined by (i) formulation/ (ii) using a CRM/ (iii) characterizing the PT item using a CRM, or (iv) using consensus value from expert labs which are not participants. In this scenario, the assigned value should be compared with robust average of participants’ results.
When Assigned value is determined by consensus value of participants, the assigned value should be compared with an independnent estimate of the assigned value such as from formulation or value reported by a competent lab which is not a participant.

## Slide 4: How to validate Assigned Value? - Scenario A
Determine Assigned Value, xpt and its std. uncertainty u(xpt) by (i) formulation/ (ii) using a CRM/ (iii) characterizing the PT item using a CRM, or (iv) using consensus value from expert labs which are not participants as stipulated in clauses 7.3, 7.4, 7.5 and 7.6 of ISO 13528:2022 respectively.
Calculate the robust average of the participants’ results x* and its std. uncertainty u(x*) using Algorithm A.
Calculate the combined std. Uncertainty (a) in Assigned Value and (b) in robust average using the formula:
udiff = Square root of [u(xpt)2 + u(x*)2]
Compare the difference between xpt and x* with 2 times the udiff
If ABS(xpt - x*) ≤ 2 x udiff : Assigned Value is validated

## Slide 5: How to validate Assigned Value? - Scenario A
Example -1 Pesticide residue in Milk |  |  |  |  |  | 
Assigned Value= |  |  |  |  | 445.3 | mg/L
EU of Assigned value= |  |  |  |  | 17.26 | mg/L
Robust average of 43 participants= |  |  |  |  | 421.58 | mg/L
Robust SD of 43 participants= |  |  |  |  | 21.47 | mg/L
SU of Assigned Value= |  |  |  |  | 8.63 | mg/L
SU of 43 participants= |  |  |  |  | 4.09 | mg/L
Combined SU= |  |  |  |  | 9.55 | mg/L
2 times Combined SU= |  |  |  |  | 19.10 | mg/L
Difference between AV and Roust Avg= |  |  |  |  | 23.72 | mg/L
AV is Validated? |  |  |  |  | NOT VALIDATED | 

## Slide 6: How to validate Assigned Value? - Scenario A
Example -2 Aflatoxin content in ground nut |  |  |  |  |  | 
Assigned Value= |  |  |  |  | 425.7 | mg/kg
EU of Assigned value= |  |  |  |  | 17.28 | mg/kg
Robust average of 19 participants= |  |  |  |  | 431.2 | mg/kg
Robust SD of 19 participants= |  |  |  |  | 14.16 | mg/kg
SU of Assigned Value= |  |  |  |  | 8.64 | mg/kg
SU of 19 participants= |  |  |  |  | 4.06 | mg/kg
Combined SU= |  |  |  |  | 9.55 | mg/kg
2 times Combined SU= |  |  |  |  | 19.09 | mg/kg
Difference between AV and Roust Avg= |  |  |  |  | 5.5 | mg/kg
AV is Validated? |  |  |  |  | VALIDATED | 

## Slide 7: How to validate Assigned Value? - Scenario B
Determined Assigned Value, xpt and its std. uncertainty u(xpt) by using consensus value from participants as stipulated in clauses 7.7 of ISO 13528:2022.
Obtain value of the PT item xref and its std. uncertainty u(xref) from an independent competent lab which is not a participant. (e.g) Homogeneity average of PT item. Instruct this lab to report EU also from which std. Uncertainty can be calculated by dividing the reported EU by k (coverage factor).
Calculate the combined std. Uncertainty (a) in Assigned Value and (b) in homogeneity average using the formula:
udiff = Square root of [u(xpt)2 + u(xref)2]
Compare the difference between xpt and xref with 2 times the udiff
If ABS(xpt - xref) ≤ 2 x udiff : Assigned Value is validated

## Slide 8: How to validate Assigned Value? - Scenario B
Example 3: Polyester content in textiles |  | 
Assigned Value from 20 participants= | 67.30 | %
SDPA= | 1.22 | %
SU of assigned Value= | 0.34 | %
Homogenity average | 66.8 | %
EU (max) reported with Homogeneity= | 1.34 | %
SU due to Homogeneity= | 0.67 | %
Combined uncerty= | 0.75 | %
2 times Combined uncerty= | 1.50 | %
Diff. Between A.V and Hom. Average= | 0.50 | %
Assigned Value validates???? | VALIDATED | 

## Slide 9: How to validate Assigned Value? - Scenario B
Example 4 : Mercury content in food |  | 
Assigned Value from 17 participants= | 152.5 | mg/kg
SDPA= | 6.48 | mg/kg
SU of assigned Value= | 1.96 | mg/kg
Homogenity average | 160.4 | mg/kg
EU (max) reported with Homogeneity= | 5.24 | mg/kg
SU due to Homogeneity= | 2.62 | mg/kg
Combined uncertainly= | 3.27 | mg/kg
2 times Combined uncertainty= | 6.55 | mg/kg
Diff. Between A.V and Hom. Average= | 7.90 | mg/kg
Assigned Value validates???? | NOT VALIDATED |